## Slide 1
S.Sureshkumar
email: skumar@ncra.tifr.res.in
GMRT Observatory, NCRA-TIFR,
Pune University Campus, Pune 411007
The Design and Performance of RF Frontend and Signal Transport System of GMRT

## Slide 2
Y shaped GMRT array is with 30 antennas of 45 m diameter parabolic dish spread over 14 km radial distance.
All antennas are connected through optical fiber to the central signal processing station located at array center.
uGMRT provide seamless frequency coverage from 120 MHz to 1450 MHz
Introduction

## Slide 3
Giant Meterwave Radio Telescope - GMRT

## Slide 4
Prime Focus Antenna Feeds of GMRT
Designed By Haumanth Rao

## Slide 5
Prime Focus Antenna Feeds of GMRT
Designed By Hanumanth Rao

## Slide 6
Prime Focus Antenna Feeds of GMRT
Designed By Hanumanth Rao

## Slide 7
Corrugated Horn
Quad-Ridged Orthomode Transducer
1000-1450 MHz Band Coverage
Prime Focus Antenna Feeds of GMRT
Designed By RRI

## Slide 8
Compact and Light Weight L band feed for GMRT
Designed By Hanumanth Rao

## Slide 9
150 MHz folded thick dipole
610/233 MHz Dual Co-axial Feed
327 MHz Cross Dipole with Sleeve
Thick dipole with Line Reflector
Other Feed designed for GMRT
Designs by G.Shankar

## Slide 10
Impedance Bandwidth
Edge Taper over feed Bandwidth
Phase Center adjustment.
Sensitivity / Deflection
Feed Performance optimization

## Slide 11
FRONT-END BOX
BYPASS
Temp
Sensor
32 Monitor Channels
POLARIZER
RFCM
Card
LNA
BroadBand
Amplifier
BPF
PostAmp
& Phase
Switch
Power
Monitor
CH1
CH2
COMMON BOX
BAND SELECTOR
BAND SELECTOR
14dB
30dB
SOLAR
ATTENUATOR
Switched
Filter Bank
Channel
Swap
Temp
Sensor
16 Control
Channels
RF ON/OFF
Noise
Gen
CW Cal
Power
Monitor
Termi
u-GMRT Front-End Block Diagram

## Slide 12
Bandwidth : 120 - 250 MHz
Gain: > 35 dB
Noise Temp. : 40 Deg. K
Return Loss: Better than -10 dB
OIP3: +20 dBm
Designed By: V. B. Bhalerao
120 – 250 MHz LNA

## Slide 14
Bandwidth : 200 - 500 MHz
Gain: 30 dB
Noise Temp. : 35 Deg. K
Return Loss: Better than -10 dB
OIP3: +18 dBm
Designed By: Anil Raut
200 – 500 MHz LNA

## Slide 15
Bandwidth : 550 - 900 MHz
Gain: 30 dB
Noise Temp. : 30 - 35 Deg. K
Return Loss: Better than -10 dB
OIP3: +18 dBm
Designed By: Anil Raut
550 – 900 MHz Low Noise Amplifier

## Slide 17
Bandwidth : 1000 - 1500 MHz
Gain: 45 dB
Noise Temp. : 32 Deg. K
Return Loss: Better than -10 dB
OIP3: +14 dBm
Designed By: Anil Raut & Sougata Chatterjee
High Gain Low Noise Amplifier

## Slide 19
Summary of RF Frontend Performance

## Slide 20
Band 2 ( 120-260 MHz ) Cascaded Bandpass Filter and Band Reject Filter for TV Transmission

## Slide 21
Designed By Imran Khan
Band 4 ( 550 – 900 MHz ) Bandpass and Switched Filter bank

## Slide 22
Designed By: Sougata Chatterjee
Band 3 ( 250 – 500 MHz ) Switched Filter Bank – A Metamaterial Design to reduce the filter size

## Slide 23
FM Radio Station ( 78 – 118 MHz ), Rejection > 50 dB
Wireless Service (150 – 175 MHz ), Rejection > 40 dB
TV (163-190 MHz), Rejection > 40 dB
TV (500 – 560 MHz), Rejection > 40 dB
Satellite Interference (230 – 280 MHz), Rejection > 40 dB
Mobile Transmission at
GSM-CDMA ( 845 – 1005 MHz ), Rejection > 70 dB
GSM 1800 MHz ( 1705 – 1940 MHz ), Rejection . 40 dB
Low Pass Filter 1650 MHz to reject out of Band signals
Designed By Ankur
Band Reject Filters

## Slide 24
Features:
    Wide Bandwidth DC- 3 GHz
Receiver with Transimpedence Amplifier
DWDM Transmission Available
High Dynamic range
Single mode fiber
Transmission distance > 40 Km
Application:
Antenna remoting - RF over fiber
Cost per Link 1600 USD
DWDM Based RF Over Fiber Optic System

## Slide 28
ƛ1
ƛ2
ƛ3
ƛ4
ƛ1
ƛ2
ƛ3
ƛ4
RFCH1
RFCH2
RFCH1
RFCH2
IF + control signals
LO Ref+ control signals
LO Ref+ control signals
IF + control signal
1550 nm
1310 nm
1550 nm
1310 nm
Universal 1.3/1.5 um
Coupler
Universal
1.3/1.5 um
Coupler
DWDM
Multiplexer
DWDM
De-Multiplexer
Network Architecture
Using DWDM Multiplexing and Bi-directional Link over Single Fiber

## Slide 29
3 Channel DWDM based Rf over Fiber Optic System
Cost per antenna 5600 USD

## Slide 31
Cost Per link : 250 USD
RF over Fiber links for Phased Array Antenna

## Slide 32
60 Element Fiber Routing

## Slide 33
Optical Specification
Wavelength: 1310 nm
Laser: Uncooled Fabry Perot Laser
Photodiode: PIN Photodiode
Optical Power Output: + 2 dBm
Optical Margin for no loss no Gain system: 3 dB
RF Specification with Pre and Post Amplifier
RF Bandwidth: 0.6 – 1.7 GHz
Input and Output Impedance: 50 Ohms
RF Gain: 6 dB
Input P1dB: > - 5 dBm
Compression Dynamic Range (CDR):130 dB/Hz
Physical Specification
Transmitter weight:116 grams
Receiver weight : 80 grams
DC Specifications
Optical Transmitter: DC 5 V / 120mA, DC-5V / 90 mA

## Slide 34
FPA with Fiber Optic Link - Test Results

## Slide 35
RF over Fiber optic link from antenna prime focus

## Slide 36
RF Frontend and Fiber Optics Team
THANK YOU